## GEX Doc#100-262
Date: |  | Signature: | Date:
Technician Name:
Instrument ID: |  | Reviewed By: | Date:
Temperature (C or F):
Calibration Standards Set ID#:
Standards Set Calibrated? (Yes/No):
Spectronic Standards Calibration Verification Tests for Genesys 20
All testing is performed with the instrument in the Transmittance mode
0% Transmittance and Stray Radiant Energy Test
| %T Result | Acceptance Criteria | Pass/Fail
0% T at 400nm: | Enter | 0
SRE at 400nm: | Enter | ≤0.45
SRE at 340nm: | Enter | ≤0.45
Wavelenght Accuracy Test
Enter Peak (nm) from Certificate: | %T Result
-5
-4
-3
-2 |  | Standard Peak Value(nm): | 0
-1 |  | Peak Found at(nm):
|  | Acceptance Criteria: | ±4nm
1
2 |  | Pass/Fail
3
4
5
Photometric Performance
Standard Value at 590nm | %T Result | Acceptance Criteria | Pass/Fail
±1.2 %T
±1.2 %T
B3 Film Peak Testing
All testing is performed with the instrument in the Absorbance mode
Test Wavelength (nm) | Absorbance Result
545
546
547
548
549
550
551
552 |  | Expected Peak(nm): | 552
553 |  | Peak Found at(nm):
554 |  | Acceptance Criteria: | ±2nm
555
556 |  | Pass/Fail
557
558
559