{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["GEX Doc#100-262"],"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"GEX Doc#100-262","index":0,"n_rows":54,"n_cols":4,"rows_kept":54,"merged":[[8,0,8,3],[9,0,9,3],[10,0,10,3],[16,0,16,3],[30,0,30,3],[36,0,36,3],[37,0,37,3]],"flags":["formulas_cached_only"]}
[]
["Date:",null,"Signature:","Date:"]
["Technician Name:"]
["Instrument ID:",null,"Reviewed By:","Date:"]
["Temperature (C or F):"]
["Calibration Standards Set ID#:"]
["Standards Set Calibrated? (Yes/No):"]
[]
["Spectronic Standards Calibration Verification Tests for Genesys 20"]
["All testing is performed with the instrument in the Transmittance mode"]
["0% Transmittance and Stray Radiant Energy Test"]
[null,"%T Result","Acceptance Criteria","Pass/Fail"]
["0% T at 400nm:","Enter",{"v":0,"d":"0.0"}]
["SRE at 400nm:","Enter","≤0.45"]
["SRE at 340nm:","Enter","≤0.45"]
[]
["Wavelenght Accuracy Test"]
["Enter Peak (nm) from Certificate:","%T Result"]
[-5]
[-4]
[-3]
[-2,null,"Standard Peak Value(nm):",{"v":0,"d":"0.0"}]
[-1,null,"Peak Found at(nm):"]
[null,null,"Acceptance Criteria:","±4nm"]
[1]
[2,null,"Pass/Fail"]
[3]
[4]
[5]
[]
["Photometric Performance"]
["Standard Value at 590nm","%T Result","Acceptance Criteria","Pass/Fail"]
[null,null,"±1.2 %T"]
[null,null,"±1.2 %T"]
[]
[]
["B3 Film Peak Testing"]
["All testing is performed with the instrument in the Absorbance mode"]
["Test Wavelength (nm)","Absorbance Result"]
[545]
[546]
[547]
[548]
[549]
[550]
[551]
[552,null,"Expected Peak(nm):",{"v":552,"d":"552.0"}]
[553,null,"Peak Found at(nm):"]
[554,null,"Acceptance Criteria:","±2nm"]
[555]
[556,null,"Pass/Fail"]
[557]
[558]
[559]